## Slide 1: Q-Exactive LC-MS

## Slide 2
TurboFlow technology
Special columns
Complex valving
Extra pump
UHPLC
Fast separation
Best separation

## Slide 3: fewer sample prep steps with TurboFlow technology

## Slide 4: Why TurboFlow columns are different: Duality
Chemistry
C18
C8
Ion Exchange
Cyclone P
Cyclone
12 different types
For targeted analyses, not profiling - yet
Size Exclusion
Turbulence within column
Velocity faster than diffusion on large molecules
Bimodal Separation on One Column
Notes: What is duality? Size exclusion chromatography and Chemistry. Based on the fact that small molecules diffuse into the pores faster than large particles. And running at high velocity the large MW diffuse slower and therefore are swept out of the column before they diffuse into the pore. Where the small molecules diffuse much faster and have the ability to go into the pores. So Turbulent Flow chromatography allows under these conditions is column technology which superior selectivity based on using 2 modes of separation rather than one. We have better selectivity over other sample prep methods.
Biomodal separation size exclusion and chemistry in one column. This is what makes TurboFlow columns so different from typical HPLC columns.

## Slide 5: Turboflow column

## Slide 6: Cleaner Plasma Extracts
A 10,000-fold dilution of . matrix
B TurboFlow Method
C Generic SPE
D Protein Precipitation
A
B
C
D
1 mm, 4-12% Bis-Tris gel
20 μL applied
Equivalent matrix amounts applied

## Slide 7: Tetracycline at 500 pg/ml with TLX vs HPLC in Pig Liver Matrix
HPLC
TLX
TLX Sample shows removal of interference compared to standard HPLC!
Courtesy of Dr. Charles Yang

## Slide 8: Leuco Malachite Green at 50pg/ml in Pig Liver HPLC vs TLX turboflow
HPLC
TLX
Aria TLX Sample shows improved S/N and removal of chemical noise!
Courtesy of Dr. Charles Yang

## Slide 9
Rat Plasma
Buffered H20
50:50 MeOH:H20
TLX System
With TurboFlow method
has less ion suppression
than other 2 methods
PPT
4:1 MeOH/H20
SPE
[J. L. Herman, Cephalon, Inc., Brandywine Parkway, PA]
Reduce Ion Suppression Effects
SPE, PPT and TurboFlow method comparison
A
B
c
Notes: Ion Suppression

## Slide 10
Ubiquitous (all species)
Concentrated (mg/ml in plasma)
Relatively late-eluting
Surfactants –“matrix effectors”
Strongly retained (ghost peaks)
Unstable –degrade to fatty acids and head groups
Ref: K. C. Van Horn, Tandem Labs, BSAT Presentation 2006
Phospholipids – ion suppression

## Slide 11: 1000 Times Less Phospholipid
106
106
103
Sample precipitation enhances extraction of endogenous phospholipids
Direct injection of plasma results in a 1000-fold reduction in phospholipid signal
A: PPT
B: direct TurboFlow method

## Slide 12
Accela 1250 + Q Exactive
R = 70,000
All compound elute in 2mins, with peak width of 0.04 mins = 9 scans under the peak
UHPLC: 60+ pesticides in 2 mins, 4s wide peaks, Full scan

## Slide 13: Instrumentation: Q Exactive
Quadrupole Mass Filter
HCD Collision Cell
S-lens
C-trap Spectrum Multiplexing
Orbitrap Enhanced FT

## Slide 14: Q-Exactive specs
Mass Resolution – up to 140,000 @ m/z 200
  Scan speed decreases with resolution
  Up to 70,000 Res. at UHPLC speeds
  Resolution decreases as mass increases
  Very high resolution for small molecule analyses
Mass accuracy < 5 ppm; usually < 2 ppm
  With external calibration
Mass range 50 – 6000 m/z
Scan speed – 12,000 amu / second
Sensitivity in ppt range
Dynamic range 1 - 5000

## Slide 15: SIM
Signal visibility is dependent on whether a signal is visible above the spectrum noise
Spectrum noise is dependent on the ratio of compound within a certain ion population
0
20
40
60
80
100
Relative Abundance
0
20
40
60
80
100
195.0876
N=248402.81
195.0877
N=20741.58
NL: 1.94E8
[150.00-2000.00]
NL: 1.12E8
[190.10-200.10]
Full MS
SIM (10amu)
S/N = 745
S/N = 5400
Lowest detected
signal/scan
250330
Lowest detected
signal/scan
28240
Chart: S/N (spectrum)
| 195.08320000000001 | 195.0838 | 195.08440000000004 | 195.08510000000001 | 195.0857 | 195.08640000000079 | 195.08700000000007 | 195.08770000000001 | 195.08830000000088 | 195.089 | 195.08959999999999 | 195.09030000000001 | 195.0909 | 195.0916
S/N (FMS) | 3.1788 | 13.2129 | 21.0615 | 66.0547 | 224.7431 | 526.3742 | 853.471 | 974.1479 | 691.1722 | 380.5662 | 123.2578 | 37.9606 | 18.2496 | 10.7579
S/N (SIM10) | 55.7154 | 92.8368 | 302.4294 | 1031.7514 | 2724.1072 | 4538.9016 | 5594.9896 | 4212.6216 | 2445.3352 | 833.317 | 249.703 | 86.2921 | 50.0377
Gain in sensitivity (7x)
Sensitivity gain 5 – 10 x with SIM mode
Caffeine

## Slide 16: Long-term Mass Accuracy with External Calibration
3° C
Realistic conditions of an average lab
Temperature variations up to 3°C peak-to-peak, up to 1°C/hour
Notes: Here has to pointed out that proven mass accuracy can be expected.
Nevertheless we spec 5 ppm for external calibration and 1 ppm for internal cal.
Mass calibration is within the specs for 24 h.
Temperature compensation is working fine.

## Slide 17: Q Exactive – functions
Function | What for? | Information
Full scan Positive or negative ions | Qualitative screening Quantitation | Accurate mass
SIM | Quantitation of knowns | Very low concentration
In-source CID | Qualitative screening Quantitation | Accurate mass with extra confirmation and structure information.
CID + MS2 | MS3 structure info. | Accurate mass with extra confirmation and structure information.
MS2 in HCD cell; All Ions Fragmentation, Data Dependent or targeted like QQQ | Qualitative screening Quantitation | Accurate mass with extra confirmation and structure information.

## Slide 18: Target or Non-Target analysis workflow
SAMPLE
Target Analysis
Non-Target Analysis
Quantitation
Target screening
?
Profiling
Fingerprinting
Authenticity
Identification
Quantitation
Q-Exactive
Triple Quad
HR/AM

## Slide 19
Metabolomics
  A comprehensive analysis of all metabolites
  A measure of the fingerprint of biochemical perturbations – pattern recognition
  Useful when you don’t know what to expect
    Hypothesis generation
Targeted Analysis
  Metabolite target analysis
    E.g. Analysis restricted to metabolites of
    an specific enzyme system that is known
    to be affected by a certain perturbation
  Metabolite profiling
    E.g. Analysis focused on a class of compounds
    associated with a particular pathway
  Only find what you are looking for
Metabolomics analyses possibilities

## Slide 20
m/z window = ± 1 Da
100
200
300
400
500
600
700
800
900
1000
0
100
[M+H]+
m/z
853.0
853.5
854.0
854.5
855.0
0
852.9720
853.4727
853.9745
854.4817
z =+2
5
>1,000,000 data points
~100,000 extracted ion peaks
Peak area ranges ~ 7 orders
Much irrelevant data
Much redundant data
High quality data from the Orbitrap allows for more precise automated data processing
Need to be able to reduce the data to chemical entities
Anatomy of a UHPLC + Orbitrap Data Set
+1
+2
+3
+4
Notes: Removing chemical noise

## Slide 21: Software for Q Exactive
Xcalibur for data acquisition & initial processing
SIEVE for handling large LC-MS data sets; differential analysis, trend evaluation; discovery
ExactFinder for searching external data bases; unknown screening & routine targeted analysis
MetWorks for metabolic samples; searches for biotransformations
Mass Frontier for structural elucidation
Proteome Discoverer for proteomics data; qualitative & quantitative
ToxID for automated screening for known compounds
Notes: Expecting Metworks shipping in July

## Slide 22
= Analyte signals
Sample
~98% of lower intensity signals are eliminated
- Solvent blank
Chemical: Background Subtraction
Notes: Removing chemical noise

## Slide 23: SIEVE: Removing Noise from Statistics
Fasted
Fed
Female
Male
Components
Loss of group separation
m/z Peaks
Increased intra group variability
Notes: Clear distinction between groups once noise is removed

## Slide 24: Q-Exactive
UHPLC – ultimate separation
TurboFlow - targeted analyses in bio matrices with little prep
Accurate mass, very stable, MS/MS
Margaret Coe
m.coe@auckland.ac.nz