UNIVERSITY OF AUCKLAND

RADIATION PROTECTION PLAN

Version 7 – February 2012

University of Auckland Radiation Protection Plan

The University of Auckland Radiation Protection Plan has been reviewed and is

Endorsed by the Occupational and Environmental Health and Safety Advisory

Committee

Yours sincerely

Professor Des Gorman BSc MBChB MD (Auckland) PhD (Sydney)

Professor of Medicine

Chair of the Vice Chancellor’s Advisory Committee on Occupational and Environmental Safety and Health

Date: 6th July 2012

CONTENTS

University of Auckland Policy	1

1.	Introduction	6

2.	Licensing, Purchase and Receipt of Radionuclide and Irradiating Equipment	7

3. 	Classification of Radioactive Sources and Irradiating Equipment	10

4. 	Operational Documentation (Unsealed Sources)	11

5.	Operational Documentation (Sealed Sources)	14

6.	Operational Documentation (Analytical X-ray Equipment)	15

7.	Operational Documentation (DEXA Scanners)	18

8.	Operational Documentation (Fluoroscopes)	20

9. 	Operational Documentation (Neutron Sources)	20

10. 	Operational Documentation (Linear Accelerators and Cobalt 60 Sources)	21

11.	Occupational Hygiene and Monitoring	23

12. 	Audits and Documentation	27

13.	Disposal	28

14.	Contingency Plans	31

15.	Use of Ionising Radiation in Clinical Research	34

Appendix 1: Activity Categories for Unsealed Sources	35

Appendix 2: Syllabus for Radiation Protection Course	36

Appendix 3: Radioactive Use Log	37

Appendix 4: Internal Audit Checklist – Unsealed sources (Exempt and A) 	38

Appendix 5: Internal Audit Checklist – Unsealed sources (Cat B and C)	40

Appendix 6: Internal Audit Checklist – Sealed Sources	43

Appendix 7: Internal Audit Checklist – Analytical X-ray Equipment	44

Appendix 8: Internal Audit Checklist – DEXA Scanner	46

Appendix 9: Internal Audit Checklist – Fluoroscopes	48

Appendix 10: Internal Audit Checklist – Neutron Sources	49

Appendix 11: Internal Audit Checklist – Irradiating Equipment	50

Appendix 12: Protocol for Personal Dosimetry	52

# UNIVERSITY OF AUCKLAND POLICY

# Policy Statement

1. The University of Auckland acknowledges that radioactive material and irradiating equipment are an important and invaluable tool in education and research within the University.

2. The purchase or acquisition of radionuclides (both sealed and unsealed) material above exempt levels and irradiating equipment by University of Auckland staff will only be undertaken by holders of a current license issued by the National Radiation Laboratory (NRL).

3. The University reaffirms the primary responsibility of license holders and staff to ensure safe use, storage, transport and disposal of radionuclides.  License holders must meet the statutory requirements of the Radiation Protection Act, Radiation Protection Regulations and Codes of Safe Practice promulgated by the National Radiation Laboratory (where these Codes are applicable as conditions of licence).

4. License holders must ensure that they are able to provide adequate supervision and instruction of staff and students using radionuclides or irradiating equipment.  To do this the License Holders must nominate all users of radionuclides and irradiating equipment and ensure that they are adequately trained in safe use.  The license holder will also ensure that risk of incurring exposure from any isotopic source or irradiating equipment is kept As Low As Reasonably Achievable (ALARA)

5. The University will assist the license holders and the NRL to monitor work with radionuclides and irradiating equipment to ensure that ALARA principals are adopted.  The University will also assist license holders by providing training on radiation protection principles to enable safe use of radionuclides and irradiating equipment.

6. The University will ensure that its responsibilities with regard to irradiating equipment as well as sealed and unsealed radio-isotopic sources are met by the establishment of:

- Internal auditing and reporting systems.

- Radiation safety training

- Systems for document retention

- Management oversight of radionuclide work

7. As line managers for their respective areas, Deans and Heads of Department are an integral part of the management of safe use of radionuclides.  The Deans, Heads of Department and Hazards and Containment Manager will work with license holders to ensure that all users of radionuclides are aware and follow statutory obligations, IAEA protocols, National Radiation Laboratory Codes of Safe Practice and the University of Auckland Radiation Protection Plan.

### Responsibilities

A. University/Faculty Administration

- In conjunction with license holders, University/Faculty Administration will provide overall management oversight on the safe use of radionuclides in their respective areas of responsibility and satisfy themselves that statutory obligations with regard to purchase, use, storage and disposal of radionuclides as well as irradiating equipment have been fulfilled.  To this end they will receive reports from the Hazards and Containment Manager.

- Will inform the Deputy Vice Chancellor (Research) of any Corrective Action Report (see section 6.2) with significant health and safety implications.

- Will direct the Hazards and Containment Manager and the University Health and Safety Coordinator to disseminate information and ensure all license holders are fully informed of all developments with regard to policy.

- Will provide appropriate courses in the safe handling of radioactive materials/irradiating equipment.

- Where license holders leave employment at the University of Auckland, the University will ensure provision is made for storage of documents relating to the statutory obligations of those license holders.  It will be the responsibility of license holders to deposit records with the Health and Safety office or the Hazards and Containment Manager.

- Faculties may support the provision of appropriate facilities for shared use of radionuclides (such as Category C laboratories).

B. Deans and Heads of Department

- In consultation with license holders, will provide management oversight on the safe use of radionuclides in their respective areas of responsibility and satisfy themselves that statutory obligations with regard to purchase, use, storage and disposal of radionuclides as well as irradiating equipment have been fulfilled.  To this end they will receive reports from the Hazards and Containment Manager.

C. University of Auckland Health and Safety Coordinator

The University of Auckland Health and Safety Coordinator is responsible for coordinating all aspects of health and safety within the University of Auckland including receipt of all accident and incident reports.  The Hazards and Containment Manager will ensure that the Health and Safety Coordinator receives copies of any reports, audits and recommendations concerning the use of radionuclides and irradiating equipment.

In the event of accidental radiation exposure, the Health and Safety Office will ensure that the individual received appropriate medical assistance and will ensure appropriate follow-up monitoring of the individual is coordinated and appropriately documented.

D. License Holders

License holders are responsible for:

- The safe use, storage and disposal of unsealed radionuclide or irradiating equipment purchased or otherwise obtained under their respective licenses.  License holders will follow the relevant license conditions (which may include the relevant NRL Code of Safe Practice).

- Ensuring that adequate records of training, purchase, transfer, use and disposal specified in Sections 4-10 of this Radiation Safety Plan are maintained and available for inspection by the Hazards and Containment Manager and the NRL.

- Ensure reporting requirements as a condition of Consent to Import are completed on time.

- Ensuring that the safe receipt, use, storage and disposal of radionuclides and irradiating equipment is in accordance with University of Auckland policy and is in compliance with the Radiation Protection Act and associated Regulations and NRL policy (which may include the NRL Safe Codes of Practice).  The guiding principal will be to ensure that all exposures are kept As Low As Reasonably Achievable (ALARA).

- In the case of work with radionuclides, ensuring that all staff and students attend training provided by the Staff and Organisational Development Unit (SODU) before commencing work with the radionuclides. Licensees will document this training.  If attendance before work has commenced  is not practical, the license holder will ensure that the new staff/student will attend the next course as soon as possible.  In such cases the license holder will be responsible for ensuring that, in the interim, the untrained staff/student is made aware of the hazards, is trained adequately to ensure safe handling of the radionuclide concerned and is given specific instructions on how to handle the radionuclide.

- Where radionuclides will be used by undergraduate students, students will be given detailed instructions in the laboratory teaching material which will include information about the hazards and steps required to ensure potential exposure is kept to a minimum.

- In the case of irradiating equipment, ensuring that all users of this equipment receive general training in radiation protection principles as well as machine specific training.  Licensees are required to ensure all training is documented.

- Ensuring that the Hazards and Containment Manager is notified of the purchase of isotopic sealed sources (above exempt levels) or irradiating equipment.  License holders must also notify the Hazards and Containment Manager of any transfers or disposal of the above radionuclide or equipment.  Licensees are also required to furnish copies of reports required as a condition of a Consents to Import  to the HCM.

- Undertaking lawful disposal or transfer of radionuclides under their control prior to leaving the University of Auckland.  This will include completing any statutory notification to the NRL and notification to the Hazards and Containment Manager.

- Ensuring a copy of the Radiation Protection Act and the Safe Code of Practice is readily available to all staff/students under their supervision.

- Ensuring adequate monitoring procedures (including wipe tests) are established and documented.

- Ensuring that adequate survey meters (where applicable) are available for monitoring work and work areas in which radionuclides are held and used.

- Ensuring that adequate safety equipment (including shielding and PPE) is available for the safe use of radionuclides.

- Ensuring that there is no transfer of radionuclides to unauthorised individuals.

- Ensuring that exposure of employees is maintained as low as reasonably achievable (ALARA).

- Distributing, collecting and collating results of personnel monitoring devices for any staff where recommended or required by the NRL Safe Code of Practice

- Ensuring that any exposure of pregnant female staff/students is kept well within the NRL guidelines for ionising radiation.

- Reporting any incidents involving accidental exposure immediately to the University Health and Safety Coordinator.  Any overexposure (as defined in Safe Codes of Practice) must be reported to the Director of the NRL within 24 hours.  License holders must inform the Health and Safety Coordinator and Head of Department at the same time.

- Ensuring that radiation emission data and/or calculated doses that staff are likely to exposed to in the course of conducting experiments is recorded and retained for inspection.

E. Users of Exempt Quantities of Radionuclide

Exempt quantities of Radionuclide are as defined in Schedule 1 of the Radiation Protection Regulations, 1982.   Use of exempt quantities will under the oversight of a current licensee.  The overseeing licensee will provide guidance.  Users of exempt quantities of radionuclide (and the principal investigator) will be responsible for:

- The safe use, storage and disposal of exempt unsealed radionuclide or irradiating equipment.  The relevant NRL Safe Code of Practice will be followed.

- Documentation of radionuclide purchase will be kept as specified in Sections 4 and 5.

- Ensuring the safe receipt, use, storage and disposal of radionuclides and irradiating equipment is in accordance with the NRL Safe Codes of Practice and the University of Auckland Radiation Protection Plan.

- Ensuring that ALARA principles are followed at all times.

- Ensuring that users attend training provided by SODU before commencing work with radionuclides.  If this is not practical, the license holder will ensure that the new staff/student will attend the course as soon as possible.  In such cases the license holder will be responsible for ensuring that, in the interim, the untrained staff/student is made aware of the hazards, is trained adequately to ensure safe handling of the radionuclide concerned and is given specific instructions on how to handle the radionuclide.

- Ensuring adequate monitoring procedures are established and documented.

- Ensuring that adequate survey meters are available for monitoring work (where appropriate) and work areas in which radionuclides are held and used.

- Ensuring that adequate safety equipment (including shielding and PPE) is available for the safe use of radionuclides.

- Ensuring that the exposure of users is maintained as low as reasonably achievable (ALARA).

- Ensuring that any exposure of pregnant staff/students is kept well within the NRL guidelines for ionising radiation.

F. All Users of Irradiating Equipment and Radionuclide Sources

- All users of Irradiating Equipment and Radionuclide Sources (both staff and students) must comply with the Radiation Protection Act and associated Regulations, license conditions (which includes the NRL Safe Codes of Practice) and this Radiation Protection Plan.  The guiding principal will be to ensure that all exposures are kept As Low As Reasonably Achievable (ALARA).

- Users must comply with the directives concerning use of Irradiating Equipment and Radionuclide Sources issued by the Licence Holder and the University.  Where exempt quantities of radionuclide are used, directives issued by the University will be followed.

G. Hazards and Containment Manager

Will work primarily in an advisory capacity with University and Faculty Administration, License Holders and Radiation Safety Officers and is responsible for:

- Providing advice on all aspects of the safe use of radionuclides and irradiating equipment in the University.

- Providing regular training classes of personnel on all aspects of radionuclide use, storage and disposal via SODU programs.

- Verifying that the purchasing/processing of transfer of sources of radionuclide/irradiating equipment meets statutory and University requirements.

- Maintaining inventory of sealed sources and handheld survey meters

- Undertaking periodic audits of all facilities and irradiating equipment.  Audits will ensure adequate documentation is in place.  The Hazards and Containment Manager will report to and make recommendations to management and license holders in the University.

- Ensuring radiation disposal facilities are maintained in a safe and secure manner for University license holders.

- Ensuring license holders and all radionuclide and irradiating machine users are kept informed of University and NRL policy directives.

- Ensuring Faculty Administrations and Directors of Institutes are kept fully informed of all reports, audits and recommendations.

1. INTRODUCTION

1.1 THE UNIVERSITY OF AUCKLAND

- The University of Auckland is responsible for delivering research-based teaching at graduate and postgraduate level.

- The University is also charged with fostering research, which is primarily publicly funded and contestable.  The results of most of this research are released into the public arena.

1.2  ADMINISTRATIVE STRUCTURE OF THE UNIVERSITY OF AUCKLAND

- The Vice Chancellor, through the Deans and Heads of Department will ensure that the University's responsibilities with regard to irradiating equipment, sealed and unsealed radionuclide sources are met by the establishment of reporting and auditing systems.

- License holders approved by the National Radiation Laboratory to import and use radio-isotopic sources and/or irradiating equipment are accountable to Heads of Department and the Dean of Faculty in which they are employed.

- As line managers for their respective areas, Heads of Department and Faculty Administration will be an integral part of the management of safe use of radionuclides and ionising radiation, ensuring that license holders are aware and follow both University policy and National Radiation Laboratory guidelines.

- The Hazards and Containment Manager in conjunction with license holders will conduct internal audits of radionuclide and ionising radiation use.  The results of these audits will be reported to the University of Auckland Health and Safety Coordinator, Institute Directors, Heads of Schools and Deans of Faculty.

2. LICENSING, PURCHASE AND RECEIPT OF RADIONUCLIDE AND IRRADIATING EQUIPMENT

2.1 USE OF RADIONUCLIDES AND IRRADIATING EQUIPMENT IN THE UNIVERSITY OF AUCKLAND

(i) 	All use of radioactive materials and irradiating equipment (except as noted in (ii) below) will be under the control of a current NRL licensee.

(ii)	Exempt sources may be stored and used without the direct control of a license but a licensee must be nominated to provide oversight.  The supervisor who has direct control over exempt sources will undertake the obligations specified in the relevant section of this Radiation Protection Plan

2.2 RECORDS OF RADIONUCLIDES AND IRRADIATING EQUIPMENT IN THE UNIVERSITY OF AUCKLAND

(i) 	The Hazards and Containment Manager (HCM) shall ensure an accurate record of:

a.) 	all current licence holders in the University of Auckland

b.)		 the location of experimental work involving radionuclides and irradiating equipment

is maintained and readily available in the event of an emergency.

(ii)	The HCM shall liaise with NRL to ensure University of Auckland records of licensees and license conditions are correct.

2.3 LICENSEE ABSENCE

- Where a licensee will be on extended leave (longer than 4 weeks and less than 3 months), that licensee must apply to the NRL to appoint a ‘Person in Charge’.  Copies of the application shall be lodged with the Hazards and Containment Manager.

- Where a licensee will be on leave longer than 3 months, another member of staff shall apply to the NRL to become a licensee for that area.

2.4  CODES OF SAFE PRACTICE AND RADIATION PROTECTION PLAN

- All staff and students working with sealed and unsealed sources of radionuclides or irradiating equipment shall observe their statutory obligations under the Radiation Protection Act and its associated Regulations and, where appropriate, the relevant NRL Safe Code of Practice.

- All license holders shall observe the requirements of this University Radiation Protection Plan.

- Staff and students handling radionuclides or irradiating equipment shall have ready access to a copy of the relevant Acts, Regulations and applicable NRL Codes of Safe Practice - this access will include web-based access

2.5 PURCHASE OR TRANSFER OF RADIOACTIVE SOURCES AND IRRADIATING EQUIPMENT

Orphan sources (i.e. those sources whose origin cannot be determined) represent a risk for the university and the attendant cost of disposal is significant.  The University also accepts responsibility for any radionuclide source or irradiating equipment in the event that a license holder fails to renew their license.  The following notification provisions are designed to reduce these risks:

- The licensee shall notify the Hazards and Containment Manager of any purchase or transfer of sealed radioactive sources or irradiating equipment.

- Purchase of all radionuclide sources shall be placed against a specific account code (775095) within University financial systems.

- The HCM shall cross reference University of Auckland purchasing records and those of suppliers to verify authorisation has been obtained.  Suppliers records will also be used to verify the accuracy of Peoplesoft and individual licensee records (see Section 4).

(iv)	Receipt of licensable sealed sources and irradiating equipment shall be notified within 10 working days of receipt of goods.  Failure to comply will be brought to the attention of the Head of Department.

2.6 NRL CONSENTS TO IMPORT RADIONUCLIDE SOURCES

- Copies of consents granted by the NRL to import radionuclide sources will be sent to the HCM.  Copies of quarterly returns and notification of receipt (required under specific consents) will be sent to the HCM.

2.7 INVENTORY RECORDS OF RADIOACTIVE SOURCES

(i)		Accurate records of purchase as well as an ongoing inventory of all sealed and unsealed sources shall be maintained by License holder or user of exempt sources as detailed in Sections 4 and 5.

2.8 RECEIPT OF UNSEALED AND SEALED RADIONUCLIDE SOURCES

- The purchaser shall immediately uplift any delivery of sealed or unsealed isotope and ensure that it is placed in a secure laboratory area.

- All packages containing unsealed radionuclide in Categories B and C activities shall be monitored to ensure that isotope has not leaked in transit.

- Any suspected contamination of outer or inner packaging materials shall be reported to the Hazards and Containment Manager immediately.

- All warning labels shall be removed or rendered unreadable before disposal of packaging.

2.9 VACATING RADIATION USE AREA

- When an area or equipment where unsealed or sealed sources of radionuclide have been used is to be returned to normal laboratory use, the facility is to be disestablished or the equipment to be scrapped, the licensee shall contact the Hazards and Containment Manager.

- The licensee shall take and document all necessary steps to ensure the area or equipment is free of contamination.  In addition to any statutory requirement to notify the NRL, a copy of the documentation shall be forwarded to the Hazards and Containment Manager and the Head of Department in the first instance.

(iii)	The above requirement will not apply to use of temporary radiation use of unsealed sources areas where activities are less than Category A.  License holders shall document all surveys taken to determine that the areas are safe to return to normal laboratory use.  Licensees shall retain the results of these surveys on file.

2.10  DECLARED PREGNANCY POLICY

- The licensee or Hazards and Containment Manager shall inform any female staff or student who chooses to declare themselves pregnant of recommended dose limits that apply to pregnant women.

(ii)	Reasonable efforts shall be taken by the licensee and female staff member (which include documented instruction) to maintain effective dose to a minimum.

3. CLASSIFICATION OF SEALED AND UNSEALED RADIONUCLIDE SOURCES AS WELL AS IRRADIATING EQUIPMENT

3.1   CLASSIFICATION OF UNSEALED SOURCES

Unsealed sources will be classified according to activities stipulated in Appendix 2 of NRL C1 and Appendix 1 of this Radiation Protection Plan.  Requirements specified in this Protection Plan for handling unsealed sources will vary according to this classification

3.2   CLASSIFICATION OF SEALED SOURCES

Sealed sources will be classified according to whether they are licensable or exempt as specified in Schedule 1 of the Radiation Protection Regulations, 1982

3.3   CLASSIFICATION OF ANALYTICAL X-RAY EQUIPMENT

X-ray emitting equipment will be categorised into one of the following categories:

- Type A equipment - Open beam XRD/XRF equipment

- Type B Equipment - XRD/XRF Equipment with interlocks and shielding where exposure is highly unlikely.

- Type C Equipment - Desktop XRD/XRF equipment or Micro CT scanners with built in shielding and interlocks so that exposure is not possible unless the equipment is opened and inbuilt shielding is removed.

- Type D Equipment - Desktop or Handheld equipment with low intensity x-ray beams, such that concerted and prolonged exposure is required.

Note that electron microscopes are not included in any of the above categories

3.4 CLINICAL X-RAY EQUIPMENT – DEXA SCANNERS

These include qCT scanners and all DEXA scanners

3.5 CLINICAL X-RAY EQUIPMENT – FLUOROSCOPES

Portable fluoroscopes will treated as a separate category

3.6 OTHER IRRADIATING EQUIPMENT

These include Cobalt 60 irradiating equipment, Linear accelerators and neutron sources.   Each item will have its own requirements.

4.  OPERATIONAL DOCUMENT (EXEMPT AND REGULATED QUANTITIES OF UNSEALED SOURCES)

License holders for unsealed sources will document the following:

4.1  Records of Radionuclide purchase, Importation or Transfer

(i)	Records of purchase of unsealed sources shall be kept in the form of an inventory of primary vials purchased.  Purchase records shall show for each primary vial of isotope:

- Radionuclide and its chemical form

- Activity

- Purchase date

- Disposal or consumption of the radionuclide

(ii)  Purchase records shall be kept up-to-date act as an inventory of primary vials of radioisotope.

- The license holder shall keep a record of any transfer of radionuclide notified to the NRL.

- The licensee shall keep a record of importation under any consent to import and furnish copies of any NRL consents and reports to the HCM (see section 2.6)

4.2  Additional Operating Procedures for the Use of the Radionuclide in

Categories B and C.

- For radionuclide use in Categories B and C (see Appendix 1), licensees shall maintain accurate records of quantities remaining after each use of radionuclides.  A usage log is provided in Appendix 3.

4.3  Record of Personnel Training

- All Staff/Students working with unsealed sources of radionuclides shall attend the University of Auckland Safe Use of Radionuclides course held by the Staff and Organisational Development Unit (SODU) (the syllabus is detailed in Appendix 2).

- Licensee/supervisor shall keep a current and accurate list of persons handling unsealed radionuclides.

- Training shall be conducted prior to handling of sealed and unsealed radionuclide in quantities qualifying as Categories B and C (see Appendix 1).

- Training should be conducted prior to commencing work with other quantities of unsealed radionuclide.  If training is not possible prior to handling radionuclide the laboratory personnel shall attend the next available training session.  In such cases, the supervisor shall ensure that detailed instructions are given to the untrained user to ensure adequate supervision.  License Holders or supervisors shall ensure that these persons handling radioisotopes attend the SODU Radiation Safety Course, have records of attendance and are supervised adequately or given specific instruction.

- After 5 years since attending initial training, persons acting under instruction or supervision who are handling quantities of sealed and unsealed radionuclide qualifying as Category B and C shall be required to complete Refresher training as outlined in Appendix 2.

4.4 Laboratory Facilities

  - Unsealed sources of radionuclide shall be handled in dedicated areas designed to ensure containment of any spill and have appropriate shielding.   These dedicated areas may be established on a temporary basis for the purposes of experimentation (e.g. may consist of a labelled tray for handling tritium).

  - Dedicated areas shall have appropriate warning signage

  - If activity of 125I or 131I handled in an open container exceeds 50 MBq, then work shall be undertaken in dedicated laboratory equipped with a fume hood. All radioiodination procedures shall be conducted in these dedicated laboratory facilities.

  - Category C and radioiodination facilities shall have a logbook  documenting use and that a scan with survey meters has been undertaken after each and every use to verify that the area is clean and not contaminated.

4.5 Monitoring

    - A system of regular monitoring contamination in the dedicated area, in and around storage locations as well as disposal areas shall be established by the licensee – wipe test monitoring may be undertaken on a Faculty, School or Departmental basis on behalf of the licensees in that sector.

    - Wipe tests shall be conducted at least on a quarterly basis for exempt sources and more regularly where high activities of isotope are handled.

    - Records of wipe tests with location of wipe tests shall be maintained by the licensee, department, school or faculty.

    - Where use of survey meters is appropriate, regular monitoring with these meters should be recorded.

    - Category C and radioiodination facilities shall have a logbook documenting use and that a scan with survey meters has been undertaken after each and every use to verify that the area is clean and not contaminated.

    - Records of monitoring shall be retained for at least 10 years.

4.6 Audit

(i) 	HCM shall conduct annual audits of all areas in which unsealed isotopes are handled.  A typical audit form is shown in Appendicies 4 and 5

4 B   USE OF UNSEALED RADIONUCLIDES IN UNDERGRADUATE TEACHING

- Use of radionuclides in undergraduate teaching shall be limited to tracer studies using exempt quantities

- Students shall be provided with adequate warning and instruction in the laboratory handbook.

- Students shall be issued with gloves and trays to ensure proper containment of radiolabeled solutions.

- Wipe tests shall be conducted as soon as possible after the completion of the section of the laboratory program involving isotopes.

5. OPERATIONAL DOCUMENTATION (SEALED SOURCES)

5.1 Inventory

License holders of sealed sources will undertake an inventory of all sealed sources under the control of License holder detailing:

- radionuclide

- activity (with reference date)

- any unique identifier (serial number)

- exempt status

- date of wipe test (license holder will conduct wipe tests for sources above exempt quantities).  This will not apply to sources where it is unsafe to perform direct wipe tests.

5.2 Wipe Tests

- Wipe test shall be conducted on all sealed sources using survey meters at a minimum.  The minimal detection activity of the survey meter will be documented for each set of tests.

- Wipe tests will be documented and results retained.  Results will be displayed as Becquerels and conversion calculation will be given.

- Any source with leakage rate of more than 0.2 kBq will be withdrawn immediately and disposed.

- This will not apply to sources where it is unsafe to perform direct wipe tests.

- Wipe tests will be performed every two years and for sources older than 30 years every 12 months.

- Records of monitoring shall be retained for at least 10 years

5.3 Training

- The License Holders shall keep an accurate and up-to-date list of all persons who use sealed sources of radionuclide that are not exempt.

- All records will be kept for a minimum of 10 years (see Section 4.1)

5.4 Audit

(i) 	HCM shall conduct annual audits of all areas in which sealed sources are handled.  A typical audit form is shown in Appendix 6

6. OPERATIONAL DOCUMENTATION (ANALYTICAL X-RAY EQUIPMENT)

License holders for analytical x-ray equipment shall document the following:

6.1 Security of Equipment

(i) 	License Holders will ensure that the irradiating equipment is secure and cannot be operated by an untrained or unsupervised person.

6.2 Local Rules - Operating Procedures for the use of the Equipment.

- License holders shall document procedures for the safe operation of a particular piece of equipment.  These shall be deemed to be Local Rules

- License holders shall ensure staff and students are aware and comply with these procedures/Local Rules.

- Specific instructions shall be given to ensure only trained license holders or suitably licensed staff undertake repair of these machines.

- All Local Rules shall detail actions in the event of an emergency.

6.3 Record of Training on Use of the Equipment

- License Holders shall keep an accurate and up-to-date list of all persons who use irradiating equipment.

- License Holders shall ensure that every staff member and students who operates equipment capable of delivering ionising radiation knows and understands Operating Procedures/Local Rules before the machine is used by that student or staff member.  In some cases these Operating Procedures may draw heavily on procedures specified by the manufacturers in the equipment handbook.

- Users are also required to have core knowledge in radiation safety.  This core knowledge can be obtained by undertaking a self-directed learning program available on the Health and Safety website in the case of x-ray equipment.  This training may be customised to each facility to ensure relevance.

- The License holders shall have a record that the staff and student displays the required competency and understanding of Local Rules to safely operate that equipment

6.4 User logs for the equipment and key holders having access to the Equipment

Logs for all irradiating equipment detailing use of equipment will be kept beside each machine.

6.5 Equipment Maintenance

Equipment maintenance shall be carried out by a person possessing a license to maintain analytical X-ray equipment (as per clause 2.4.1 NRL C17), the responsibility for the maintenance licensee will cease once machine has been repaired, safely re-commissioned and documentation of verification completed (see below).  Safe Use of the machine then becomes the responsibility of the Use licensee.

6.6  Record of Equipment Maintenance

  - License holders will maintain record of all maintenance performed on the x-ray equipment.

  - License holders will pay particular attention to the ability of maintenance personnel to bypass safety interlocks and remove shielding.  In such cases, license holders will document verification that interlocks have been reinstated, shielding has been reinstated safely and dose rates are within acceptable limits.

  - Maintenance licenses shall provide a written statement to record verification that interlocks are functioning, shielding has been properly reinstated and dose rates are within acceptable limits.

6.7   Monitoring Protocols

(i)	Irradiation equipment leakage and scatter will be measured at intervals specified in Section 6.8 or after any alteration to the machine.

(ii) All records will be kept for a minimum of 10 years.

6.8 Variations for Different Types of XRD/XRF Equipment

      - Type A equipment - Open beam XRD/XRF equipment

All of the above requirements apply with leakage/scatter monitoring undertaken on a 3 monthly basis

      - Type B Equipment - XRD/XRF Equipment with interlocks and shielding where exposure is highly unlikely

All of the above requirements will apply with annual leakage/scatter measurements

      - Type C Equipment - Desktop XRD/XRF equipment with built in shielding and interlocks so that exposure is not possible unless machine is opened

Local rules may be solely drawn from manufacturers instructions but must include emergency procedures. Radiation Protection Training may consist of a modified form of web-based safety training.  Annual leakage/scatter measurements to be undertaken

      - Type D Equipment - Desktop equipment with low energy beams, such that concerted and prolonged exposure is required

Local rules and training may be solely drawn from manufacturers instructions but must include emergency procedures.  No leakage measurements are required.

6.9 Audit

(i) 	HCM shall conduct annual audits of all areas in which analytical x-ray equipment is used.  A typical audit forms are shown in Appendix 7

7. OPERATIONAL DOCUMENTATION FOR DEXA SCANNERS

7.1 Documentation

Where DEXA scanners are used on patients, License holders for DEXA Scanners shall ensure there are procedures to document:

- the name of the patient (or identifier) and referring clinician for clinical studies

- that all clinical trials performed on the DEXA scanner have appropriate ethical approval and that all subjects have signed the approved consent form for the scanning procedures.

- low dose mode is used for children and all scans performed on a children are clinically appropriate .

7.2  Warning signage and Consent

Warning signage on entrance to areas containing DEXA scanners shall be present Patients shall be informed of the dose that they are likely to incur.  It is recommended that dose is equated to a meaningful measure to a lay person (such as number of hours on an international flight).

7.3 Scanning Records and QA records

License holders shall ensure that:

- A software system or a machine use log is maintained so that it is possible to:

- Cross-reference scan data with the patients name (or identifier).

- Determine average number of scans per day

- QA is run each day before use.

- Regular phantoms are performed and meet the established machine normal range and these phantom scans are reviewed regularly by the license holder to ensure safe and accurate performance. If any issues arise with the performance of the scanner, this is to be reported to a qualified service technician and rectified before further use.

- QA and phantom data is retained and retrievable.

7.4 Record of Training on Use of the DEXA Scanners

- License Holders shall ensure that every person who operates DEXA scanners understands the Operating Procedures/Local Rules before the scanner is used.

- License Holders shall keep an accurate and up-to-date list of all persons who are specifically authorised by the license holder to use DEXA Scanners (i.e. that the licensee is satisfied that the person displays the required competency and understanding of operating procedures to safely operate that equipment).

- Users are also required to have core knowledge in radiation safety.  This core knowledge can be obtained by undertaking a self-directed DEXA programme available on the Health and Safety website.  This training may be customised to each facility to ensure relevance.

- Where clinical studies are undertaken, it is preferable that operators undertake registered DEXA training (such as offered by the ANZBMS) which will meet the requirements of (iii) above.

7.5 Equipment Maintenance

(i) 	Scanner maintenance shall only be carried out by a person possessing a license to maintain DEXA x-ray equipment.  The licensee is notified once the machine has been repaired and safely re-commissioned and the responsibility for ongoing safe usage of the DEXA machine is returned to the licensee.

- Maintenance license holders shall document verification to ensure that if safety features have been bypassed or removed, these features have been reinstated, and dose rates are within acceptable limits before use.

- The license holder shall retain copies of maintenance reports for 10 years

7.6  Scatter Measurements

- Operating procedures shall ensure the operator is seated at the machine recommended distance or shielded if space does not allow adequate distance between the machine and the operator.

- Where shielding is used, licensees shall retain records of verification of shielding efficiency.

- All records will be kept for a minimum of 10 years.

7.7 Audit

(i) 	HCM shall conduct annual audits of all areas in which DEXA scanners are used.  A typical audit forms are shown in Appendix 8

8. OPERATIONAL DOCUMENTATION FOR FLUOROSCOPES

Where fluoroscopes are used in animal studies, the license holders for fluoroscopes shall document the following:

- The equipment is only operated only by licensee or a suitably qualified radiographer

- Log of use

- Annual certification of Machine by certified Medical physicist

- Annual certification of lead aprons, gloves and thyroid screens provide shielding.

- All others present in the room must have written instructions to stand at least 1 meter from operating table while x-ray machine in operation

Room where the fluoroscope is to be used must have documented shielding or environmental dose rates shall be recorded every 2 years

All records shall be kept for a minimum of 10 years.

Audit

(i) 	HCM shall conduct annual audits of all areas in which fluoroscopes are used.  A typical audit forms are shown in Appendix 9

9. OPERATIONAL DOCUMENTATION FOR NEUTRON SOURCES

The license holders for the neutron shall document the following:

- The equipment is only operated only by licensee or a suitably qualified nurse

- Log of use

- Verification of dose to patient and operator

- Environmental survey at dose at various locations in the facility

All records shall be kept for a minimum of 10 years.

Audit

(i) 	HCM shall conduct annual audits of all areas in which neutron sources are used.  A typical audit forms are shown in Appendix 10

10. OPERATIONAL DOCUMENTATION FOR LINEAR ACCELERATORS and COBALT 60 IRRADIATORS

10.1 Security of Equipment and Authorised Access

License holders shall:

- take reasonable steps to keep access to facilities with Cobalt 60 sources and linear accelerators to a minimum.

- keep an accurate register of key holders and reconcile those with access to training records.

- ensure that every staff member/student who operates Linear accelarators and Cobalt 60 irradiators understand Operating Procedures/Local Rules before the machine is used by that student or staff member.

- ensure key holders who access to the facilities but do not operate the equipment are given instructions concerning the hazards and what to do in an emergency.

- ensure emergency instructions and up-to-date contact list is posted in the log book or in a prominent place in the Facility.

10.2 Training

- License Holders shall keep an accurate and up-to-date list of all persons who have access and who use irradiating equipment.

- License Holders shall ensure that every staff member/student who operates Linear accelarators and Cobalt 60 irradiators understand Operating Procedures/Local Rules before the machine is used by that student or staff member.

- Local Rules will include Emergency protocols

- Users are also required to have core knowledge in radiation safety.  This core knowledge can be obtained by undertaking a self-directed learning program as part of the training or by attending the SODU radiation protection safety program.

- The License holder shall have a record that the staff and student displays the required competency and understanding of Local Rules to safely operate that equipment with signoff by the licensee.

10.3 Personal Dosimetry Records

License holders shall ensure:

- all users wear dose monitoring devices (such as film badges) and have these read at regular intervals (as per Appendix 12).

- records of dose received are kept in a secure place and retained for at least 10 years.

- missing records are accounted for.

- all readings from microprocessor dosimeters are recorded in a logbook.

- microprocessor dosimeters are calibrated annually.

10.4 Use of Equipment

License holders shall ensure accurate records of use are kept.

10.5 Maintenance

License holders shall ensure the following are undertaken and documented:

  - all interlocks and warning devices are tested on a 6 monthly basis

  - 6 monthly environmental survey with positions for readings marked on floor plans

  - 6 Monthly testing of any warning alarms.

10.6 Audit

(i) 	HCM shall conduct annual audits of areas in which linear accelerators and Cobalt 60 irradiators are used.  A typical audit forms are shown in Appendix 11

Summary of documentation

License holders for Linear accelarators and Cobalt irradiators shall document the following:

- Log of use

- Maintenance log (this may be part of Log of Use)

- 6 monthly test of interlocks and warning lights

- 6 monthly environmental survey with positions for readings marked on floor plans

- 6 Monthly testing of any warning alarms

- Annual check of any pocket dosimeter used in conjunction with film badges

- Annual test for leakage where Cobalt 60 sources

- List of users and their training

- List of users and personal dosimetry checks (film badges) as per procol in Appendix 9

All records will be kept for a minimum of 10 years.

11. OCCUPATION HYGIENE AND MONITORING

11.1 PERSONNEL MONITORING

- Laboratory personnel handling 125-I or 131-I with activities greater than 17 MBq (0.5 mCi) in a volatile form (i.e. performing radio-iodinations) shall have a thyroid scan once a week or after each occasion radioactive iodine is used, which ever is less frequent.  A survey monitor fitted with a scintillation detector may be used for monitoring within one week of the manipulation.  Readings shall be recorded in the local logbooks.  If any thyroid uptake is detected, a calibrated thyroid scan using a suitable neck phantom and used in a reproducible geometry.

- Calibrated thyroid scans shall be performed by Nuclear Physics Department, Auckland Hospital.  Copies of reports shall be sent to licensee and be retained by the HCM.

- Staff are handling 32-P in quantities greater than 34 MBq (1.0 mCi) in any single experiment shall be monitored using finger TLDs.

- Staff or students using 60Co sources shall be monitored using film badge or microprocessor dosimeter.  Dosimeters will be calibrated every year.  Procedures for issuing and reviewing dosimeters are detailed in Appendix 12.

- Staff handling more than1.5 GBq (45 mCi) 3-H or 14-C should submit a urine sample within 48 hours of using the radionuclide to ensure ingestion of radionuclide has not occurred.  Once permission has be obtained from staff submitting samples, results of any monitoring shall be retained by the University Health and Safety Office ensuring that generally accepted standards pertaining to retention of information and respect for individual privacy are observed.

11.2 DOSE MONITORING

- The Hazards and Containment Manager shall monitor doses received in the course of typical experiments involving sealed and unsealed sources.  Results shall be published on the University Health and Safety Website.

- Dose levels shall be measured around Category A x-ray diffraction equipment when it is possible for users to have access to the immediate vicinity of the x-ray beam.  The measurements shall be undertaken every 3 months and shall be documented

11.3 RADIATION DOSE LIMITS

Notwithstanding the statutory requirement to maintain dose limits within those limits specified in ICRP 60, the University will adopt ALARA principles.  Limits for persons exposed to radiation as a normal condition of employment should be maintained at 1/3 limits specified in ICRP 60.

| Effective dose to body | 6mSv |
| --- | --- |
| Equivalent dose to skin | 150mSv |
| Equivalent dose to eye | 50mSv |
| Equivalent dose to hands or feet | 150mSv |
| Equivalent dose to abdomen of pregnant women | 0.6mSv |

11.4 CONTAMINATION MONITORING

  - Wipe tests shall be undertaken at least on a 6 monthly basis to monitor use of unsealed low energy beta emitters.  Records shall be forwarded to the Hazards and Containment Manager or may be retained in the laboratory for inspection.

  - Survey Monitors will be used to monitor work area for contamination where higher energy beta and gamma emitting radionuclides are used.  Work areas shall be monitored after each use.

11.5  INVENTORY OF MONITORING AND SURVEY EQUIPMENT.

- The Hazards and Containment Manager shall maintain an accurate register of monitoring and survey equipment along with calibration status.  Monitoring equipment shall include radiation area monitors near Cobalt 60 sources and other irradiating equipment.

- Calibration of survey equipment shall be determined using a check source of known activity.  Probes will be placed as close as possible to check source containing either beta or gamma source to obtain maximum possible reading and to minimise variation contingent on distance and orientation of check source.

- Where possible, detection efficiency readings obtained will be compared to average efficiencies for the same meter.  Efficiencies that are lower than 30% from average will indicate further investigation is required.  Survey equipment that fails to meet acceptable standards shall be withdrawn from service for repair or disposed

- These records shall be kept with the Hazards and Containment Manager.

- University dosimetry equipment used to verify local dosimetry readings shall be calibrated every 4 years.

11.6  ENVIRONMENTAL SCANNING.

- Areas surrounding high activity gamma sources (ie. > 1 Ci) or neutron sources shall conduct environmental scans every 2 years.  A plan showing position of readings shall be part of the documentation. Readings shall be taken with source in exposed and shuttered position.   Documentation will be retained with each instrument.

- Leakage and scatter measurements will be taken every 12 months employing a phantom (DEXA Scanners) or a sample of high density material (micro CT)

12. AUDITS AND DOCUMENTATION

12.1 INTERNAL AUDIT

The over-riding consideration in conducting any audit shall involve facilitation with licensees to ensure safe working practices are adopted.

The following audits shall be carried out by the HCM:

  - Audit of documentation for each license holder

  - Environmental surveys of exposure levels and tests for contamination.

  - User training verification.

  - Comparison of survey meters.

  - Audit of work areas – including signage.

Note: Examples of audit documents are given in Appendices 4-11.

- The frequency of audit shall be at least once every 12 months but will ultimately depend on use/purchase of radionuclide and levels of risk involved.

- As a result of audit, the University Radiation Protection Plan shall be reviewed to ensure currency and adequacy.

- License Holders and Head of Department shall be advised of the outcome of the audit.

12.2 REQUESTS FOR CORRECTIVE ACTIONS

- Where internal or NRL audits result in corrective action requests, the Head of Department shall be informed of these requests.

- The Deputy Vice Chancellor (Research) shall be advised of any corrective action request where significant hazard is present.

12.3 AUDIT REPORTS

- A copy of the audit reports shall be retained by the Hazards and Containment Manager for inspection.

- A summary of the reports will be presented to the Deans of the relevant Faculty and OESHAC.

12. 4 TRAINING RECORDS

- In addition to licensee obligations for training outlined in Sections 4-10, the Hazards and Containment Manager shall ensure the following:

a) 	Undergraduate students using radionuclides are issued with adequate instruction.

b)	An awareness program for maintenance staff is established.

12.5 RETENTION OF RECORDS

- Licensees shall retain all records relating to sealed and unsealed radionuclides or irradiating apparatus (including dosimetry) for 10 years.

- Where a license holder relinquishes their license or leaves the University, all records relating to the radionuclides or irradiating apparatus shall be forwarded to the Hazards and Containment Manager.

- It shall be the responsibility of the license holder to ensure that such records are deposited with the Hazards and Containment Manager.

12.6 RADIATION PROTECTION PLAN

- The Radiation Protection Plan shall be issued with date and version number.

- A copy shall be lodged with the Health and Safety Office.

- The Hazards and Containment Manager shall be responsible for ensuring currency of the Plan.

- The Hazards and Containment Manager shall ensure Faculties and license holders are aware of any changes or amendments to the Radiation Protection Plan

13. DISPOSAL

13.1 DISPOSAL OF SEALED SOURCES AND IRRADIATING EQUIPMENT

- Records of disposal or transfer of ownership of irradiating equipment and equipment containing licensable sealed sources of radionuclides shall be forwarded to the National Radiation Laboratory.

- Records of disposal or transfer shall include any source that may be present in scintillation counters.

- A copy of the transfer or transfer shall be sent to the Hazards and Containment Manager.  Disposal of equipment shall follow procedures outlined in relevant Codes of Safe Practice.

13.2 DISPOSAL OF UNSEALED RADIOISOTOPIC SOURCES

Wherever possible, decay to activities less than 0.01x activities specified as limits for disposal will be used a method of disposal

Solid Waste

- Where radionuclide has a half-life of less than 90 days, waste shall be allowed to decay to levels specified by the relevant National Radiation Laboratory Safe Code of Practice.  To enable reasonably accurate calculation of the time allowed to decay isotopic waste, licensees shall keep good approximation of isotopic activities disposed.  No allowance will be made for decay prior to storage of waste to decay.

- All warning signs of containers and overpacks shall be removed before disposal.

- Designated decay facilities shall be used for storage for the purposes of decay.

- Detachable labels shall be used for waste containers stored in decay facilities.

- For radionuclides with half life longer than 90 days, waste shall be disposed of in accordance with Appendix 5 of National Radiation Laboratory Safe Code of Practice.

- License holders are strongly advised to allow a margin of safety of at least 5 in their calculations (i.e. they should halve the limits used in Appendix 5 of NRL C1)

Liquid Waste

- Where radionuclide has a half-life of less than 90 days, waste shall be allowed to decay to activities 10 times less than specified for sewer by the National Radiation Laboratory Safe Code of Practice NRL C1 before disposal to sink.

- Designated decay facilities shall be used for storage where decay is being employed as a method of disposal.

- Detachable labels shall be used for waste containers stored in decay facilities.

- To enable reasonably accurate calculation of the time allowed to decay isotopic waste, licensees shall keep good approximation of isotopic activities disposed.  No allowance will be made for decay prior to storage of waste to decay.

- For unsealed liquid radionuclides with half-life longer than 90 days, waste shall be disposed of in accordance with Appendix 5 of the National Radiation Laboratory Safe Code of Practice NRLC1, with the exception that allowable levels for disposal at sink shall be those specified for sewer by the Safe Code of Practise.

- License holders are strongly advised to allow a margin of safety of at least 10 in their calculations.

- Disposal to sewer shall be via labelled sink. The sink and associated exposed pipe fittings should be routinely monitored.

Syringes, needles, glass pasteur pipettes, razor blades

Sharp material will be packed and stored in approved plastic pails specifically used for sharp rubbish and allowed to decay wherever possible.

Contaminated Radionuclide Waste with hazardous chemical or biological properties

    - Waste contaminated with radionuclides which also possesses hazardous chemical or biological properties cannot be treated with single protocol.

    - All radionuclide waste with hazardous chemical or biological properties will be notified to the Hazards and Containment Manager to ensure protocols that deal with multiple properties are drafted.

    - Wherever possible, biodegradable liquid scintillation cocktails with high flashpoints will be employed to facilitate disposal.

13.3 RECORDS OF DISPOSAL

Records of disposal shall be kept for all radionuclides in Category C Laboratories and retained for 10 years.

13.4 RADIONUCLIDE DECAY FACILITIES

The Hazards and Containment Manager shall inspect all facilities set aside for radionuclide decay every 6 months to ensure facilities have adequate security, adequate signage and storage procedures are observed.  The HCM shall conduct contamination monitoring every 12 months.

14. CONTINGENCY PLANS

14.1 ACCIDENTS/INCIDENTS

- The Hazards and Containment Manager and the Health and Safety Coordinator shall be informed immediately in the event of:

    - Injury related to work with radionuclides/irradiating equipment

    - Exposure to ionising radiation resulting an equivalent dose of 0.6 mSv.

    - Any unaccounted loss of radionuclides.

    - Theft of radionuclides/irradiation equipment

    - Sabotage of irradiation equipment

- Although it is anticipated that reports should be forwarded via license holders, license holders shall be notified immediately.

- Accident/incident reports shall be reported on the University report form and forwarded to Health and Safety Office.

- Where the accident/incident involves significant potential exposures (as detailed in section 14.3) the license holder shall inform the National Radiation Laboratory within 24 hours.  The license holder will also inform the Faculty Manager (via the Head of Department) and the following will be undertaken:

a.	An internal investigation will commence immediately.

b.	Corrective action will be undertaken as soon as possible.

c.	In the event of a breach in security access control or locks would be changed within 24 hours.

- Records of the incident, internal investigation and corrective actions shall be retained by the license holder, the Hazards and Containment Manager and the Health and Safety Office.

14.2 ACCIDENT/INCIDENT INVESTIGATION

- All accidents and incidents involving radionuclides or irradiating equipment shall be reported to the University of Auckland Health and Safety Office in the manner prescribed for all accidents and incidents.

- Both Head of Department and License holder shall be informed prior to reporting to the H&S Office.

- The H & S Office shall advise the Hazards and Containment Manager of any accidents or incidents involving radionuclides or irradiating equipment in order that the Hazards and Containment Manager can conduct an investigation.

14.3 ACTION LEVELS

  - Where the accident results in a potential exposure level greater than 10% of the dose detailed below or results in ingestion of radionuclide of more than one-twentieth of the ALI (annual limit of intake), license holders shall immediately inform the Health and Safety Office and their Faculty Manager (via Head of Department).

  - Where the accident results in a potential exposure level greater than detailed below or results in ingestion of radionuclide of more than three-tenths of the ALI (annual limit of intake), license holders shall immediately inform the National Radiation Laboratory, Health and Safety Office and their Faculty Manager (via Head of Department).

Note: the ICRP equivalent dose limits for pregnant women are lower than other staff members.

| Effective dose to body | 6mSv |
| --- | --- |
| Equivalent dose to skin | 150mSv |
| Equivalent dose to eye | 50mSv |
| Equivalent dose to hands or feet | 150mSv |
| Equivalent dose to abdomen of pregnant women | 0.6mSv |

14.4 SPILL

HCM shall be notified of any spill to ensure adequate decontamination and arrange follow up wipe tests if necessary.

14.5 PROCEDURES IN THE EVENT OF AN EMERGENCY

a. Unsealed sources at Category C Level

Procedures in the event of an emergency shall kept in each category C laboratory

b. Irradiating Equipment

Standard Operating Procedures shall be kept with each item of equipment detailing procedures to be adopted in the event of an emergency and nominating the person responsible for ensuring the equipment is turned off before evacuation.

14.6 EMERGENCY PLANS

Emergency Plans will identify location on floor plans of unsealed and sources of radionuclide and irradiating equipment in the University of Auckland.

14.7 RISK MANAGEMENT PLAN

- The Hazards and Containment Manager will assess the risks posed by use and storage of radionuclides and irradiating apparatus for the University of Auckland.

- The Hazards and Containment Manager will develop and update a contingency plan to manage these risks in the event of an emergency.

15. USE OF IONISING RADIATION IN CLINICAL RESEARCH

15.1 USE OF RADIATION IN CLINICAL RESEARCH

  - License holders shall comply with the relevant NRL Code of Safe Practice for clinical practice.

  - License holders shall take all reasonable efforts to ensure recommended calibration procedures are undertaken at intervals specified by the manufacturer.  Results of these calibration procedures shall be retained on file by the licensee.

  - Patients shall be informed of the typical dose of ionising radiation they will receive in the course of the procedure.  Comparison with typical doses received in other clinical or dental x-ray procedures may be useful to ensure patients are adequately informed.

  - Warning signs shall be posted warning pregnant women of the presence of x-rays or ionising radiation.

  - Every effort shall be made to ensure paediatric doses are used (to ensure lowest effective doses are employed) where children are part of clinical studies

  - License holders shall employ a licensed radiographer where directed to do so by the NRL or by a NRL Code of Practice

  - Hazards and Containment Manager shall receive a copy of all correspondence with the NRL.

  - APPENDIX 1:  Activity Categories

| Radionuclide | O | Category A | Category B | Category C |
| --- | --- | --- | --- | --- |
| H-3 | < 1 GBq < 30 mCi) | 1 GBq - 10 GBq 30 - 300 mCi | 10 Gbq - 100 GBq 300mCi - 3 Ci | 1 TBq 3 - 30 Ci |
| C-14 | < 10 MBq <300 uCi | 10 MBq - 1 GBq 300 uCi - 30 mCi | 1 - 10 GBq 30 - 300 mCi | 10 - 100 GBq 300 mCi - 3Ci |
| P-32 | < 30 kBq <1 uCi | 30 kBq - 100 MBq 1 uCi - 3mCi | 1 GBq 3 - 30 mCi | 10 GBq 30 - 300 mCi |
| P-33 | < 30 MBq <1 mCi | 30 - 100 MBq 1 - 3mCi | 100MBq - 1 GBq 3- 30 mCi | 1 - 10 GBq 30- 300 mCi |
| S-35 | < 30 MBq <1 mCi | 10 - 100 MBq 1 - 3 mCi | 100 MBq - 1 GBq 3 - 30 mCi | 10 GBq 30 - 300 mCi |
| I-125 | < 3 MBq <100 uCi | 3 - 10 MBq 100 - 300 uCi | 10 - 100 MBq 300uCi - 3 mCi | 100 MBq - 1 GBq 3 - 30 mCi |

Activities Requiring Documented Usage

| Radionuclide |  |
| --- | --- |
| H-3 | > 10 GBq (> 300 mCi) |
| C-14 | > 1 GBq (> 30 mCi) |
| P-32 | > 100 MBq (> 3 mCi) |
| P-33 | > 100 MBq (> 3 mCi) |
| S-35 | > 100 MBq (> 3 mCi) |
| I-125 | > 10 MBq (> 300 uCi) |

APPENDIX 2: SYLLABUS FOR RADIATION SAFETY COURSE

The training programme consists of one three hour workshop sponsored by SODU which covers the following topics:

- BACKGROUND THEORY

  - Nuclear decay

  - Interaction with matter and living tissue

  - Dose and its measurement

  - ICRP threshold limits and their application to radiation safety

- SAFE HANDLING OF RADIONUCLIDES

  - Basic principles of ALARA

  - Basic principles of Radiation Safety

  - Legal Obligations and University Policy

  - Transport

  - Sealed Sources

  - Disposal

  - Protocols for emergencies.

REFRESHER TRAINING

    - After 5 years since attending initial training, persons acting under instruction or supervision who are handling quantities of sealed and unsealed radionuclide qualifying as Category B and C shall be required to attend a one hour Advanced Safe Handling of Radionuclides course held by the Staff and Organisational Development Unit.

    - Persons who are eligible for retraining (as per (i) above) will be exempt from attendance at the Advanced Safe Handling of Radionuclides course if they can demonstrate a good understanding of Radiation Protection.  One method of demonstrating such an understanding is to achieve a minimum 70% score on a MCQ test which is part of the University Radiation Licensee Training Program

Appendix 3:

UNIVERSITY OF AUCKLAND

Radioactive Usage Log

(for Category B and C activities only)

Location:

Radionuclide and chemical form:

Date:

| Date | Name | Activity removed | Activity Remaining |
| --- | --- | --- | --- |

#### APPENDIX 4:  INTERNAL AUDIT CHECKLIST – UNSEALED RADIONUCLIDES

#### (Exempt and Category A levels)

#### License Holder:

#### Department and Room Number:

#### Isotope Used:				Activities:

#### Documentation

#### License required		 Code of Practice readily available		

#### Records of purchase		  List of users				

#### Users   – all attended CPD/SODU course?					

#### Substandard Items:

#### Monitoring program

#### Method					Frequency

#### Area included in Wipe Test program		Date of last check:

#### Laboratory Facilities

#### Prominent warning signs on doors		Surfaces impervious		

#### Washbasins/soap/towels		Fume Hood available		

#### Hot Area

#### Clearly defined				Adequate signage			

#### Free of Clutter					Adequate containment – trays etc	

#### User Log?					Hot sink labelled			

#### Contamination Check (32P; 125I only)					

#### Substandard items:

#### Shielding

#### Correct shielding 				Adequate Shielding			

#### Survey instrument

#### Correct type					Battery charged			

#### Included in UoA survey		

#### Storage Areas

#### Fridges/Storage areas labelled		Containers labelled			

#### Activities present match records		

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### Follow up Audit (where necessary)

#### To check items:

#### Performed by: _________________________  Date: __________

#### APPENDIX 5:  INTERNAL AUDIT CHECKLIST – UNSEALED RADIONUCLIDES

#### (Category B and C levels)

#### License Holder:

#### Department:

#### Room Number:

#### Isotope Used:				Activities:

#### Frequency of use:

#### Documentation

#### Current License		 Code of Practice readily available		

#### Records of purchase		Records of use 					

#### List of users			

#### Substandard Items:

#### Monitoring program

#### Method					Frequency

#### Documentation

#### Substandard Items:

#### Laboratory Facilities

#### Prominent warning signs on doors		Surfaces impervious		

#### Washbasins/soap/towels		Fume Hood available		

#### Refrigerator labelled	

#### Hot Area

#### Clearly defined				Adequate signage			

#### Free of Clutter					Adequate containment – trays etc	

#### User Log?					Hot sink labelled			

#### Correct shielding 				Adequate Shielding		

#### Substandard items:

#### Survey instrument

#### Correct type					Battery charged			

#### Included in UoA survey		

#### Storage Areas

#### Fridges/Storage areas labelled		Containers labelled			

#### Activities present match records		

#### Contamination Check (32P; 125I only)

#### Area clean				  Storage Area clean			

#### Training and OHS

#### Users   – all attended CPD/SODU course?			

#### Emergency protocols? (Class C only)			

#### Thyroid monitoring records			

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### Follow up Audit (where necessary)

#### To check items:

#### Performed by: _________________________  Date: __________

#### APPENDIX 6:  INTERNAL AUDIT CHECKLIST – SEALED RADIONUCLIDE SOURCES (Excluding fixed sources inside Irradiators)

#### License Holder:

#### Department:						Room Number:

#### Documentation

#### Inventory Present			 Inventory correct/matches license	

#### Log-out/Login Procedure		List of Users			

#### Users trained				

#### Wipe Tests

#### Date of Last Wipe test:

#### Minimum Detectable Activity documented 

#### Results given in Becquerel 	Conversion calculations shown	

Sources with leakage greater than 0.2 kBq?  Y/N

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### Follow up Audit (where necessary)

#### To check items:

Performed by: _________________________  Date: __________

#### APPENDIX 7: INTERNAL AUDIT CHECKLIST – ANALYTICAL X-RAY EQUIPMENT  (XRD and XRF)

#### License Holder:

#### Department:

#### Room Number:

#### Make:					Model:

#### Category of X-ray equipment (A-D):

#### Documentation

#### Current License			Log of use					

#### Local Rules/Procedures		

#### Procedures to Prevent Unauthorised Access

#### Prominent warning signs on doors					

#### Access to control panel controlled when not in use			

#### Log of key holders							

#### Equipment key actuated?						

#### Software password							

#### Door locked when equipment unattended				

#### Equipment Audit

#### Equipment labelled with warnings				

#### Equipment warning lights functioning				

#### Interlocks functioning				

#### Dosimetry readings

#### Leakage and scatter survey undertaken				Date:

#### NB: Not required for Category D devices

#### Work Practices

#### Name of XRD Operator:

#### Is that person only operator?				Yes/No

#### If not , is there a log of training			Yes/No

#### List of all users			All users trained		

#### Training method

#### Do Local Rules incorporate Emergency Procedures			

#### Maintenance

#### Maintenance undertaken and logged					

#### Service agent for machine	Yes/No  Name:

#### Maintenance staff are licensed to repair					

#### Sign off procedures before recommissioning					

#### Evidence of perimeter dose rate before recommissioning			

#### Emergency:

#### Is there an Emergency protocol prominently displayed?			

#### Is there a Person responsible for shutdown?					

#### Location of emergency shutdown switch known or is

#### the switch is clearly labelled?						

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### APPENDIX 8: DEXA SCANNER

#### License Holder:

#### Department: 				Room Number:

#### Make:					Model:

#### Documentation

#### Current License				

#### Current Ethics approvals available		

#### System to log machine use			

#### List of all users and their training record	

#### Operating Procedures				

#### QA and Phantom record system		

#### Evidence of review of phantom scans	

#### Are children likely to be scanned?	Is there awareness of use of paediatric made?

#### Procedures

#### Prominent warning signs on doors					

#### Warning signs for pregnant women					

#### Door locked when equipment unattended				

#### Dosimetry readings

### Date of Scatter survey undertaken if appropriate:

#### Room layout or screening appropriate for machine		

#### Maintenance

#### Regular maintenance undertaken and logged					

#### Licensed Service agent for machine	Yes/No

#### Name:

#### Written documentation of safe re-instatement				

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### APPENDIX 9: FLUOROSCOPE

#### License Holder:

#### Department: 				Room Number:

#### Make:					Model:

#### Documentation

#### Only operated only by licensee or a suitably qualified radiographer	

#### Operator dose monitored by

#### Log of use								

#### Written instructions for others present				

Annual certification of Fluoroscope:

Undertaken by:

Date:

Annual certification of lead aprons, gloves and thyroid screens 	

Adequate shielding in walls						

Adequate warning signage						

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### APPENDIX 10: NEUTRON SOURCE

#### License Holder:

#### Department: 				Room Number:

#### Source:

#### Documentation

#### Only operated only by licensee or a suitably qualified Nurse		

#### Log of use									

Verification of dose to patient and operator					

Environmental survey at dose at various locations in the facility		

Adequate warning signage							

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### APPENDIX 11:  INTERNAL AUDIT CHECKLIST –IRRADIATING EQUIPMENT

#### (Linear accelerator and Cobalt 60 sources)

#### License Holder:

#### Department:						Room Number:

#### Type of equipment					Activity of Source:

#### Make:							Model:

#### Documentation

#### Current License			 Code of Practice readily available	

#### Maintenance Log Present		 Log of use				

#### Laboratory Facilities

#### Prominent warning signs on doors	 Room locked when not in use		

#### Log of key holders			

#### Equipment

#### Equipment labelling	

#### Equipment warning lights and interlocks functioning	

#### Date of last check:

#### Dosimetry readings

#### Perimeter and typical dosimetry checks last taken:

#### Environmental survey last taken:

#### Pocket dosimeters in use 		Date of last calibration

#### NRL Film Badges 			Film badges cover all users 

#### Frequency of renewal:

#### Length of time records retained:			Location of file:

#### Work Practices

#### User instructions and/or SoP				SoP  for maintenance		

#### Training

#### Up-to-date		

#### Covers Radiation Protection principles	

#### Emergency:

#### Accidents/incidents		Emergency protocol		

#### Person responsible for shutdown		Emergency contacts		

#### Comments:

#### Audit performed by: _________________________  Date: __________

#### Follow up Audit (where necessary)

#### To check items:

Performed by: _________________________  Date: __________

APPENDIX 12:  UNIVERSITY OF AUCKLAND

Protocol for Personal Dosimetry

- Personal dosimeters shall be issued by licensee responsible for irradiating apparatus or high activity radionuclide source.

- The licensee shall keep a record all staff and students issued with film badges and TLDs

- Personal dosimeters may be film badges or thermo-luminescence dosimeters (TLDs) issued by the National Radiation Laboratory or microprocessor dosimeters that record cumulative dose at each occasion the dosimeter is used.

- Licensee shall issue instructions on correct use and correct position that dosimetry device is to be worn.

- Licensee shall ensure all film badges or TLD are renewed at least every 3 months.

- In the case of microprocessor dosimeters, the licensee shall ensure all readings are recorded in a logbook.

- Microprocessor dosimeters shall be calibrated annually.

- The licensee shall retain all records of film badge or TLD results for all staff and students issued with film badges or TLDs.

- The licensee shall review results and ensure that dose exceeding 100microSieverts is reported to the HCM immediately.

- The licensee shall record all lost or non-returned film badges and TLDs.  The licensee shall ensure a new monitoring device is obtained prior to the individual continuing work with radionuclide or apparatus.

- The HCM shall verify compliance with the above procedure at time of internal audit.